PURA HYGROMETER

-120C to -40C Range

4-20 mA Sensor Output

INSTALLATION, OPERATION AND

MAINTENANCE MANUAL

##### Table of Contents

1.	INTRODUCTION	4

1.1	General	4

1.2	PURA Hygrometer	4

1.3	Ceramic Technology Sensor	4

1.4	Calibration	4

1.5	Manufacturing Quality	5

2.	PREPARATION	6

2.1	Installing the Monitor	6

2.2	Electrical Connections	7

2.2.1	Electrical Power	8

2.2.2	Setpoint (Alarm) Relays	8

2.2.3	Analog Output	8

2.2.4	Digital Interface	8

2.3	Sensor Cable	9

2.4	PURA Sensor Installation	9

3.	INSTRUMENT SETUP	11

3.1	Setup Security Feature	11

3.2	Selecting the Engineering Units	12

3.4	Analog Output Calibration	15

3.5	Analog Output Scaling	15

3.6	Display Brightness Adjustment	16

3.7	Digital Communications	16

4.	PRESSURE COMPENSATION	18

4.1	Using an Optional Pressure Transducer	18

4.2	Using a Fixed Pressure Input in Single Channel Mode	19

4.3	Optional Pressure Transducer Connection	19

5.	OPERATION	21

5.1	Sampling Hints	21

5.2	Response Characteristics	22

5.3	Which Gases to Measure?	22

6.	TECHNICAL SPECIFICATIONS	24

6.1	Monitor	24

6.2	Sensor	25

7.	MAINTENANCE	26

APPENDIX 1	27

Register settings accessible by digital communication	27

APPENDIX 2	29

Setup Codes	29

APPENDIX 3	32

Error Messages	32

# 1.	INTRODUCTION

## 1.1	General

Kahn Instruments manufactures and distributes a wide range of on-line and portable instruments for the determination of moisture content in air and gases. In addition, the company manufactures humidity calibration systems for industrial and scientific users.

## 1.2	PURA Hygrometer

The PURA hygrometer is a continuous, on-line instrument for the measurement of absolute moisture content in gas. It is designed to operate under a wide range of conditions and provide for the monitoring and/or control of moisture in gases. The instrument consists of two component parts: monitor and sensor, each of which is individually calibrated to a single standard allowing combinations of sensor and monitor units to be totally interchangeable.

The standard instrument covers the ranges -120 C to -40 C dewpoint (-184 F to

-40 F dewpoint), 0-9999 ppmv as well as 0-99.99 ppbv. Selection of the displayed hygrometric unit is factory set but may be easily changed by the user. Dual alarm relay contacts are provided which are user configurable both in terms of setpoint and operating mode (i.e. NC or NO). Optionally, one or two additional relays can be fitted with Form A type contacts (see section 2.2.2 for details). Current output is standard and factory set at 4 - 20 mA. Optional outputs are 0 - 20mA or 0 - 10 VDC.

The monitor has a pressure input channel to accept the signal from any standard 2-wire pressure transmitter. Pressure can be scaled and displayed to appropriate engineering units. The pressure signal can then be used to provide pressure compensation on the primary channel of ppmv or ppbv.

## 1.3	Ceramic Technology Sensor

The PURA hygrometer utilizes the Kahn ceramic technology moisture sensor.  This sensor is manufactured from metalized ceramics using thin and thick film technologies to produce a sensor, which is virtually chemically inert with inherently fast response, high calibration stability and high resilience to corrosive environments. The PURA hygrometer displays moisture content by converting the analog 4-20 mA signal from the intelligent micro-controller based ceramic moisture sensor.

## 1.4	Calibration

Calibration is performed using precision dewpoint generators and transfer standard optical hygrometers, which have been calibrated directly at the National Institute of Standards and Technology (NIST).

As with all precision measurement equipment, calibration integrity depends on regular verification. Kahn Instruments suggest that the instrument be returned to the Kahn Calibration Laboratory at least once a year for a re-calibration.  Alternatively an exchange sensor can be ordered and simply installed by the user on-site to avoid costly downtime.  In either case, calibration is inexpensive and ensures continued accurate operation of the instrument.

## 1.5	Manufacturing Quality

Your hygrometer should reach you in perfect working condition.   Rigorous procedures at every stage of production ensure that the materials of construction, manufacturing, calibration and final test procedures meet the requirements laid down by our Quality System.

# 2.	PREPARATION

The PURA hygrometer is supplied with a full accessory package. Please check that all the following standard components are present in the packing box:

1)	Monitor - PURA hygrometer

- Moisture Sensor

- Sensor cable assembly

- Power cable

Any additional items will be shown on the packing list.

## 2.1	Installing the Monitor

The monitor requires an operating environment of 0 C to 50 C, 0 to 90% RH.

The monitor case is designed for panel mounting. However, it can be used as a bench mounted device without any special preparation.

For panel mounting, a suitable cutout should be made in the panel to be used. A rectangular cutout of 92 x 45 mm (DIN 1/8) is necessary (see Figure 1. Panel Cutout). Optional NEMA 4 cover provides additional front panel protection for harsh environments.

Figure 1. Panel Cutout

NOTE:	A minimum depth behind the panel of 5.5” is required for the monitor, its connectors and wiring (see Figure 2. Case Dimensions).

The monitor should be inserted into the panel cutout from the front and secured using the two brackets with securing screws on the sides of the front case.

All electrical connections through the rear panel are made via plug and socket terminals to facilitate easy removal of the monitor when in service.

## 2.2	Electrical Connections

When the monitor is securely fastened in position, the electrical connections can be made to the back panel.

Figure 2. Case Dimensions 3 shows the connections for power, dewpoint sensor, alarm relays and analog output. These details are also repeated on the instrument itself. The exact factory setup of the monitor unit can also be obtained by reference to the connection details label that is applied to the monitor case.

Figure 3. Instrument Wiring Diagram

### 2.2.1	Electrical Power

The power supply to the monitor may be one of the following:

- 86-265 VAC 50/60Hz - factory default setting

- 10-48VDC or 18-36 VAC - optional

The power supply voltage is indicated on the label located on the monitor.

As soon as power is connected at the rear of the monitor, the display will register a reading. The monitor is provided for continuous operation and does not feature a power ON/OFF switch.

### 2.2.2	Setpoint (Alarm) Relays

SP1 and SP2 are Form C relay contacts and can be connected as Normally Open (NO) or Normally Closed (NC). Contacts are rated 240 VAC/10 A or 24 VDC/8 A (non-inductive load).

Optional additional relays SP3 and Open Sensor (OS) have Form A relay contacts and are Normally Open only.  Contacts are rated 240 VAC/5 A or 30 VDC/ 5 A (non-inductive load).

Configuration of setpoints is covered in section 3.3.

### 2.2.3	Analog Output

The monitor as standard provides a 4 to 20 mA (optionally: 0 to 20mA or 0-10VDC) linear output over the entire operating range of the instrument in the selected engineering units.  The output can be digitally scaled to cover a partial range of the measurement parameter. See Section 3.5 for details.

The maximum load resistance that can be connected between pin 16(+) and pin 17(-) for a CURRENT output is 500 Ohms (No minimum).

The minimum load that can be connected to pin 16(+) and pin 17(-) for a VOLTAGE output is 5 kOhms (No maximum).

### 2.2.4	Digital Interface

The monitor can be equipped with an optional RS232 or RS485 digital interface. A two meter RS232 cable, terminated with a 9 way ‘D’ socket, is available on request.

## 2.3	Sensor Cable

The sensor cable should be connected to the monitor by its three wires as shown on Figure 3. The standard cable length is 6 feet. If additional cable length is required, it can be supplied on request.

If the customer prefers to use its cable, then it is important that the cable meet the following specification: three 7/0.25mm (22 AWG.) stranded, copper conductors, insulated with polypropylene, twisted together with a common copper shield wire, wrapped in aluminized polyester tape and enclosed in a PVC outer sheath.

## 2.4	PURA Sensor Installation

If you purchased the PURA Premium, PURA OEM or PURA sensor models with the optional cleaning to oxygen standards, it has been assembled and packaged within a class 100 clean-room environment. To maintain this level of cleanliness the packaging should only be breached within the same or cleaner environment.

The effective operation of the PURA, in a flowing gas environment, relies on the sensor being installed directly into the gas stream or by having a fully representative gas sample directed over the sensor measurement surface. Where possible avoid installing the sensor in a “dead” or unswept volume.

The use of Swagelok® retained gasket assemblies, containing silver plated, stainless steel ¼” VCR gaskets is recommended when connecting the PURA into a gas line. The distance between the inlet and outlet gas connection ports is set at a pitch of 120 mm.

Install the sealing gasket onto the VCR connections on either the PURA or the connecting gas lines. Ensure that the PURA is offered into the gas line with reference to the gas flow direction and the inlet port as indicated on the PURA body. Tighten the female nut firmly finger tight. While holding the PURA stationary with a spanner, tighten the gas line nut 1/8 (one eighth) of a turn using a second spanner. Repeat this operation on the remaining gas connection port.

CAUTION: Over tightening the nuts can cause permanent damage to the seals and seatings.

Fig 4.  PURA Premium & PURA OEM Models.

Finally, connect the sensor cable to the sensor using the four-pin connector already provided on the cable. The instrument is now ready for configuration and use.

# 3.	INSTRUMENT SETUP

NOTE: When the instrument is first powered up the display may show a zero value for about 1 second, followed by a flashing “OPEN” message for approximately 5 seconds, before showing a dewpoint value. This is normal and does not indicate a problem with the instrument.

There are two levels of operation, USER and ADVANCED.

User allows:

- Changing of the setpoint values;

- Monitoring of the min/max values.

Advanced programming allows setting of the following options:

- advanced setpoint programming;

- analog output calibration and scaling;

- setting the engineering units.

To access the advanced programming press and release SETUP and  simultaneously to scroll through the menus.

## 3.1	Setup Security Feature

To prevent unauthorized access to the internal settings, the monitor is equipped with two DIP-switches located on the display board of the monitor, which can be accessed underneath of the faceplate (refer to Figure 5. Location of the LOCKOUT Switches.). The PROGRAM LOCKOUT switch SW2 is responsible for enabling or disabling the code programming mode. The SETPOINT LOCKOUT switch SW1 enables or disables the setpoint programming. OFF position of either switch will enable that mode, ON position will disable that mode.

Figure 5. Location of the LOCKOUT Switches.

## 3.2	Selecting the Engineering Units

Enter the PROGRAM UNLOCK mode as described in Section 3.1 Setup Security Feature, and press SETUP and  to scroll through the menus.

The selection of moisture measurement engineering units is achieved by entering different CODE 4 and CODE 7 as shown below: (See Appendix 2, CODE 4 & CODE 7 for full details).

| Display | CODE 4 | CODE 7 |
| --- | --- | --- |
| Dewpoint in C | 000 | 000 |
| Dewpoint in F | 010 | 000 |
| PPMV | 000 | 100 |
| PPBV | 010 | 104 |

To set the range and resolution for ppmv and ppbv, set digit 3 of CODE 7 to:

0 	1 – 9999 for ppmv

1	0.1 – 999.9 for ppmv

2	0.01 – 99.99 for ppmv

3	0.001 – 9.999 for ppmv

4	0.01 – 99.99 for ppbv

For example, by setting CODE 7 to 102, the primary display shows ppmv with a resolution of 0.01 to a maximum of 99.99, above which, message oVEr will be displayed. The secondary display shows dewpoint in degrees Celsius.

When configured to display PPMV or PPBV the monitor is also capable of displaying dewpoint in degrees C or F. The indication in degrees is accessible by pressing the  button. Press the SETUP button to go back to the primary variable PPMV or PPBV. Dewpoint in degrees C or F selection is determined by the second digit of the CODE 4, see table above.

3.3	Changing the Setpoint Values

Before amending the setpoint values, make sure that the monitor is in the SETPOINT UNLOCK mode, as described in 3.1 Setup Security Feature.

Press SETUP and  simultaneously to gain access to the setpoint codes.

The setpoints of the relays are set by setting SP_n (where n = 1 to 4 and represents SP1 to 4) to the required level. The units of the setpoints change according to the source of the setpoints, i.e. if the indicator is displaying ppmv, then the setpoints are set in ppmv.

For the relays to be energized above the setpoint SPC_n = 0xx.

For the relays to be energized below the setpoint SPC_n = 1xx.

The source of relay setpoints are set by setting digit 2 of SPC_n  i.e. xnx as shown below:

0	Displayed value  (DEFAULT)

1	Dewpoint

2	Pressure

3	Sensor connection fault

For example; if you wish the relay to trip at a certain dewpoint value while the indicator is displaying ppmv, then SPC_n = x1x, or if the relay is to trip if a failure occurs with the sensor, then SPC_n = x4x.

The secondary relay function is set by digit 3 of SPC_n  i.e. xxn as shown below:

0	No function   (DEFAULT)

1	Relay latched

2	De-energize relay

3	No function

4	Relay de-energized for sensor connection fault

5	Relay energized for sensor connection fault

6	Relay toggles at 1Hz for sensor connection fault

7	Setup for hysteresis, delay type and make/break delay times

For example, if you wish the relay to trip above the setpoint using the displayed value, but to remain latched, then SPC_n = 001. It will then remain latched until the supply to the indicator is removed or SPC_n = 002.

The function of the relay annunciators is set by digit 1 of CODE 1 (i.e. nxx), as shown below:

0	LED annunciators  always OFF

1	LED annunciators ON when relays are de-energized

2	LED annunciators ON when relays are energized (DEFAULT)

Hysteresis, Make/Break Delay & Delay Type

Associated with each setpoint is a Hysteresis Value, Make delay time, Break delay time and a delay type. To gain access to these parameters, set SPC_n  = xx7. To scroll the features press  or .

The hysteresis value is a value that the relay trips above and below the nominal setpoint. For example, if SP1 is set to -80.0 and the HYST is set to 2.0 and the relay is set to energize above the setpoint, then the relay will energize when the setpoint source rises above -78.0 and is de-energized when the setpoint source falls below -82.0.

The Make delay is the delay between the setpoint being reached and the relay energizing. For example, if M_d = 0.02.30 then the relay will be energized 2 minutes and thirty seconds after the trip point has been reached.

The Break delay is the delay between the relay being energized and it de-energizing. For example, if b_d = 0.01.12 then  when the relay is energized it will remain energized for 1 minute and 12 seconds , after which it will be de-energized.

The maximum make and break delay time is 9 hours, 6 minutes and 6 seconds in increments of 1 second.

The setpoint can have four different delay types: NorM (normal), rEPt (repeat), 1Shot, and PuLSE.

If deLAY = NorM then the relay will function normally with the inclusion of the time delays.

If dELAY = rEPt then the make and break delays will repeat continually until the setpoint source goes to a level to deactivate the setpoint. For example, if the make delay = 10 sec & break delay = 5 sec, then after the trip point is reached, the relay will energize after 10secs and remain energized for 5secs after which it will de-energize for 10 sec and then energize for 5 sec etc.

If dELAY = PuLSE then if using the make and break times mentioned above, the relay will energize after 10secs, de-energize after 5secs and remain de-energized.

If dELAY = 1Shot then if using the make and break times mentioned above, the relay will energize after 10secs and remain energized.

## 3.4	Analog Output Calibration

The analog output is calibrated at the factory. To verify the 4-20 mA analog output calibration, perform the following:

- Connect milliammeter in series with 250 Ohms resistor between terminals

16 (+) and 17 (-) of the monitor

- Set CAL=071. The monitor will display message CAL_L flashing with XXXX, where XXXX is an internal digital variable. Using  or  buttons adjust variable XXXX to read 4.000 mA on the milliammeter.

- Press SETUP button. The monitor will display message CAL_h flashing with YYYY, where YYYY is an internal digital variable. Using  or  buttons adjust variable YYYY to read 20.00 mA on the milliammeter.

- Press SETUP button, then restore the previous value of the CAL code (default 052). Press SETUP button several times to exit the programming mode.

To verify the 0-10 VDC analog output calibration, perform the following:

- Connect voltmeter between terminals 16 (+) and 17 (-) of the monitor

- Set CAL=071. The monitor will display message CAL_L flashing with XXXX, where XXXX is an internal digital variable. Using  or  buttons adjust variable XXXX to read 0.000 VDC on the voltmeter.

- Press SETUP button. The monitor will display message CAL_h flashing with YYYY, where YYYY is an internal digital variable. Using  or  buttons adjust variable YYYY to read 10.000 VDC on the voltmeter.

- Press SETUP button, then restore the previous value of the CAL code (default 052). Press SETUP button several times to exit the programming mode.

## 3.5	Analog Output Scaling

Enter the PROGRAM UNLOCK mode as described in 3.1, press SETUP then , then SETUP again. Now use  and  to scroll through to the CAL menu.

The analog output can be scaled by setting CAL to 061 and setting the zero value to the required output at 4mA (or 0V) and the full scale value to the output required at 20mA (or 10V).

For example, if the output is required to go from -110.0C dewpoint at 4mA to -50C dewpoint at 20mA then set Zero to -110.0 and F.S. to -50.0

To exit the CAL menu, set CAL to 052 and use  to scroll through the menus.

## 3.6	Display Brightness Adjustment

To adjust the display brightness:

- Press SETUP and  buttons simultaneously, the display toggles between [bri] and [5], where 5 is default setting;

- Adjust the display brightness required (from 0 to 7) by using  or  buttons, then press SETUP several times to exit the programming mode.

## 3.7	Digital Communications

The monitor can be supplied with an optional ASCII RS232 or RS485 serial interface.

The communication settings can be found by setting CAL to 100, then pressing the SETUP button once and then  and  to change the baud rate. Baud rates are 300, 600, 1200, 2400, 4800, 9600, 19,200 and 57.6K. Press the SETUP button again to advance to the parity bit and then  and  to change the parity bit. Parity settings are odd, even or off. Press the SETUP button again to advance to the address settings. Address settings are 0 to 255. Note, address 0 is not a valid RS485 address. Press the SETUP button again to return to CAL.

The communications uses the following default protocol:

Baud rate 	9600

Data bits	8

Parity		none

Flow control	Xon/Xoff

The message to read and write to the monitor follows the following format:

1) Start Character

"s" or "S" for the start character (must be first character in string).

2) Meter Address

An ASCII number from "0" to "255" for the meter address. If the character following the start character is not an ASCII number, then address 0 is assumed. All meters respond to address 0.

3) Read/Write Command

The next character must be an ASCII "R" or "r" for read, or an ASCII "W" or "w" for write. Any other character will abort the operation.

4) Register Address

The register address for the read/write operation is specified next. It can be either an ASCII number from "0" to "255" or registers 1 - 18 can be accessed by entering and ASCII letter from "A" to "R" (or "a" to "r", not case sensitive). If the address character is omitted in a read command, the meter will always respond with the data value currently on the display. (The register address must be specified for a write command).

5) Separator Character

After the register address in a write command, the next character must be something other than an ASCII number. This is used to separate the register address from the data value. It can be a space or a "," or any other character except a "$" or a "*".

6) Data Value

After the separator character, the data value is sent . It must be an ASCII number in the range of "-32766" to "32766".

7) Message Terminator

The last character in the message is the message terminator and this must be either a "$" or a "*". If the "$" is used as a terminator, a minimum delay of 50mS is inserted before a reply is sent. If the "*" is used as a terminator, a minimum delay of 2mS is inserted before a reply is sent. (The "$" and "*" characters must not appear anywhere else in the message string).

Examples:

SR$			Read display value, 50mS delay, all meters respond.

s15r$			Read display value, 50mS delay, meter address 15 responds.

Sr130*		Read Code 1 setting, 2mS delay, all meters respond.

s2wb-10000$	Write -10000 to the display register of meter address 2, 50mS delay.

SW65 200$		Write 200 to the SP1 Hysteresis register, 50mS delay, all meters respond.

s10w148,7*		Change brightness to 7 on meter address 10, 2mS delay.

sr6$			Reads SP1

sw6-100$		Sets SP1 to –10.0

sr5$			Reads pressure value

A full list of the register that can be read or written to can be found in Appendix 1.

# 4.	PRESSURE COMPENSATION

Enter the PROGRAM UNLOCK mode as described in 3.1.

The monitor has the ability to measure pressure in order to provide a pressure compensated value for ppmv or ppbv.

## 4.1	Using an Optional Pressure Transducer

In order to enable the use of a pressure transducer, set CODE 4 to 107 (dual channel mode), by pressing SETUP and  to scroll through the menus.

If you wish to display Dewpoint while measuring pressure, then set CODE 4 to 107 to show Dewpoint in C or 117 in F.

To display pressure in PSIG set CODE 6 to 000, or set CODE 6 to 100 to display in BARG.

Manual Pressure Input Calibration

The pressure input channel needs to be calibrated in order to match the range of the pressure transducer. This is achieved by setting CAL to 012 and entering values for offset OFF_2 and scale SCA_2.

SCA_2 = 0.0062 per 100 PSIG

OFF_2 = pressure range – ((20,000 x (pressure range/1000))/16)

For example, for a pressure transducer with a range of 0 to 1000 PSIG:

SCA_2	= 0.0062 x 10 = 0.0620

OFF_2 = 1000 – ((20,000 x (1000/1000))/16) = 1000 – 1250 = -250

If you are using a pressure transducer in BARG, then you have to convert the value to PSIG by multiplying it by 14.5.

For a pressure transducer with a range of 0 to 100 BARG:

100 BARG = 1450 PSIG

SCA_2 = 0.0062 x 14.5 = 0.0899

OFF_2 = 1450 – ((20,000 x (1450/1000))/16 = 1450 – 1813 = -363

To exit the CAL menu, set CAL to 000 and press  to scroll through the menus.

Automatic Pressure Input Calibration

Alternatively, the pressure input can be calibrated using a 2-wire transducer simulator with 4 and 20mA output. To do this, connect the simulator between pin 6 (+ve) and pin 5 (-ve).

Set CAL to 022 and press SETUP. The monitor will then display ZEro and flash 0. Set the simulator output to 4.0mA and press SETUP to set 0 at 4mA.

Press SETUP and the monitor will display SPAn and flash the full scale pressure value. Set the simulator output to 20.0mA and use the  &  buttons to set the required pressure range. Then press SETUP to set the required pressure for 20.0mA.

Set CAL to 000 and press  to leave the menus.

## 4.2	Using a Fixed Pressure Input in Single Channel Mode

In order to display pressure compensated values of ppmv and ppbv without the use of a pressure transducer, then this can be achieved by entering the pressure value manually.

Enter the PROGRAM UNLOCK mode. Use SETUP and  to scroll through the menus and select code 7.

Select the required pressure compensated value, i.e. CODE 7 = 1xx for ppmv, 4xx for ppbv, and CAL = 052.

Exit the menus and exit the PROGRAM UNLOCK MODE. Press SETUP for two seconds, and the display will show the set pressure. To change the pressure, use  and  and then press SETUP.

## 4.3	Optional Pressure Transducer Connection

The monitor provides excitation voltage (24VDC @ 20mA) for an auxiliary 2-wire transmitter used in the dual channel configuration. Connect (+) of the transmitter to pin 6 of the monitor, (-) of the transmitter to pin 5 of the monitor. Refer to Figure 6. Wiring Diagram for details.

Press  button to monitor the pressure value. Press  button to monitor the dewpoint value. Press SETUP button to monitor PPMV or PPBV.

Figure 6. Wiring Diagram for Pressure Sensor

# 5.	OPERATION

Operation of the PURA hygrometer is very simple assuming that the necessary precautions are taken to protect the sensor from damage.  Statistical information indicates that the vast majority of failures are caused either by incorrect sampling methods, sampling positions or inadequate protection against dangerous substances.

## 5.1	Sampling Hints

The ceramic sensor is designed to operate in a flowing gas stream.  The sensor can be mounted directly into a flowing gas stream in a duct or pipe, provided there is no form of contamination within the pipe/duct, which will damage the sensor. A sensor sampling block is available which allows a small sample of process gas to be diverted past the sensor before returning to the main gas stream or being bled off to atmosphere.

The general rules to be adhered to when arranging a sampling system are as follows:

Be Sure the Sample is Representative of the Gas Under Test

The sample point should be as close to the critical measurement point as possible. For example, in a glove box application, mount the sensor at the exit of the glove box, not at the gas entry point.

Minimize Dead Space in Sample Lines

Try to avoid too many tee pieces or unnecessary tubing.  Where possible, construct a sampling system specifically for the job. Do not use tubing previously installed for another application. Dead space in sample lines causes moisture entrapment points, increased system response times and measurement errors as the trapped moisture is released into passing sample gas causing an increase in partial vapor pressure.

Remove Any Particulate Matter or Oil from the Gas Sample

Particulates at high velocity can damage the sensor.  Similarly, at low velocity, they may "blind" the sensor and reduce its response speed.  If particulates such as degraded desiccant or pipe scale and rust are likely to be present in the sample gas, use a particulate in-line filter.  Kahn technical sales personnel will be pleased to give advice.

Use High Quality Sample Tube and Fittings

We recommend that, wherever possible, stainless steel tubing and fittings be used.  This is particularly important at low dewpoints since other metals have hygroscopic characteristics and adsorb moisture on the tubing walls, slowing down response and, in extreme circumstances, giving false readings.  For temporary applications or where stainless steel tubing is not feasible, use high quality thick-walled PTFE tubing, which exhibits similar characteristics to stainless steel.

Always use the shortest run of tubing possible between two points.  Use the smallest diameter tubing possible to reduce response time, but take care not to induce pressure differentials by aiming for too high a flow rate through small diameter tubing. A sampling flow between 1 and 5 liters per minute (2-10 SCFH) will be satisfactory for the ceramic sensor to operate correctly.

## 5.2	Response Characteristics

Many factors will affect the response speed of the PURA hygrometer.  However, your task will be made simpler if, like the vast majority of applications, your PURA hygrometer is being used to detect an increase in moisture content of your sample gas.  It is much easier for the sensor to detect an increase in moisture content than to respond to a decrease.  This is similar in principle to the way it is easy to make a sponge wet, but more difficult to dry it out afterwards.

However, since response characteristics from dry to wet are orders of magnitude faster than from wet to dry, the effective response speed of the system to leaks, ingress of moisture, etc. is very fast.  In real life applications such as dryer monitoring or glove box monitoring the sensor will respond in a few seconds to an increase in moisture.

In practice, the time taken to dry down a sensor from ambient conditions to the operational dewpoint level of the process will normally be shorter than the time taken to dry down the process itself.  Therefore, when the sensor is installed into the system prior to system start-up, there is normally no time lag before representative test results are obtained.

If a new sensor is installed into an operational system, for example a gas dryer exit pipe via a sampling line with particulate filter, then typically fifteen to thirty minutes should be allowed for the tubing, filter and sensor to reach equilibrium with the sample gas passing through.

## 5.3	Which Gases to Measure?

The PURA hygrometer, by nature of its design, is suitable for measurement of the moisture content of a wide variety of gases.  In general, if the gas (in conjunction with water vapor) is not corrosive to ceramics or base metals then it will be suitable for measurement by the PURA hygrometer.  However, gases containing entrained solids or hydrocarbon mists should be filtered upstream of the ceramic sensor. High purity gases will not be contaminated by the ceramic sensor.  It has no components that are likely to outgas (epoxy, plastics, etc.) and, therefore, is safe for use in critical semiconductor and fiber optic applications.

High pressure samples may be measured up to a maximum of 5000 psig.  Make sure that the correct fittings are used and that a bonded seal is used to seal the radial face of the sensor/sensor block seal.

Similarly, the sensor is suitable for measuring samples below atmospheric pressure. Care should be taken at joints and pipe fittings that no leak exists which could allow ambient air to contaminate the sample gas. If there is doubt, a standard leak test procedure should be employed (helium or vacuum seal).

# 6.	TECHNICAL SPECIFICATIONS

## 6.1	Monitor

Display:				5 digit LED

Measurement Ranges:			-120 C to -40 C dewpoint, -184 F to -40 F dewpoint, 0 to 9999 ppmv,

0 to 99.99 ppbv user selectable.

Pressure input channel:		4-20mA input scaleable to any engineering units.

Electronic accuracy:		Dewpoint ±0.5ºC, ±1.0ºF PPMV range ±1% of reading, secondary channel ±0.5% FS.

Outputs:				4-20mA as standard or 0-20mA optional (max load 500 ohms) or 0-10V optional (min load 5 kOhms), spanned to full calibration range, scaleable by user.

Alarm Relays:			Two relays SP1 and SP2 are standard. Control action and setpoint are user programmable.  Form C contacts rated 10 A, 240 VAC or 8 A 24 VDC non-inductive load.

Relays SP3 and SP4 (OS) are optional. Control action and setpoint are user programmable.  Form A contacts rated 5 A, 240 VAC or 5 A, 30 VDC non-inductive load.

Operating environment:		0 to +50 C, 0 to 90%RH.

Power supply:			115 VAC, 50/60 Hz or 220 AC 50/60 Hz standard.

Optionally: 18-36 VAC or 9-60VDC.

Power consumption:		Max 10 Watts.

Power connection:			6 feet, 3 wire.

Weight:				1.32 lbs.

Sensor cable:			Three 7/0.25mm (22 AWG.) stranded, copper conductors, insulated with polypropylene, twisted together with a common copper shield wire, wrapped in aluminized polyester tape and enclosed in a PVC outer sheath. Fitted with sensor connector and terminations for monitor. Maximum length 800 meters.

Environmental protection:	IP54 / NEMA 12. Optional protection cover to IP66 / NEMA 4.

## 6.2	Sensor

Type:					Kahn PURA Sensor.

Calibration range:			-100 to -40 C dewpoint.

Measurement range:		-120º to -40º C dewpoint.

Interchangeability:			Fully interchangeable sensors.

Dewpoint Accuracy:		±1.0 C from -40 to -60 C

±2.0 C from -61 to -100 C

±4.0 C estimated from -101 to -120 C

(Or equivalent for other ranges).

Operating temperature:		-40C to +60C.

Gas wetted surfaces:		316 Cold Drawn Stainless Steel, 0.25 micron electro-polished.

Gas connection ports:		PURA Premium & PURA OEM: ¼” VCR fixed male ports.

Installation profile:			PURA Premium & PURA OEM: 120mm between inlet and outlet sealing face.

Temperature coefficients:	Temperature compensated from -20C to +40C

Operating vacuum / pressure:	10-9 torr vacuum to 5000 psig.

Flow rate:				Recommended 1-5 liters per minute (2-10 SCFH).

Traceable certification:		-75 to +20C dewpoint traceable to National Institute of Standards and Technology.

For dewpoints < -75C direct reference to a fundamental chilled mirror hygrometer.

Environmental protection:	IP65.

Weight:				0.33 lbs.

Fault conditions:

Overrange dewpoint:	20 mA output, factory programmed.

Underrange dewpoint:	4 mA output, factory programmed.

Sensor Failure:		23 mA	output, factory programmed.

# 7.	MAINTENANCE

Routine maintenance of the PURA hygrometer is confined to regular re-calibrations.  This work can only be done by exposure of the moisture sensor to sample gases of known moisture content. Calibration services traceable to the National Institute of Standards and Technology (NIST) are provided by Kahn Instruments.

The frequency of re-calibrations required in order to maintain the performance of the PURA hygrometer is primarily dependent on the composition of the gas to which the sensor is exposed: i.e. content of liquid and particulate contaminants, corrosive elements, etc.  In most applications, annual re-calibration ensures that the stated accuracy of the PURA hygrometer is maintained.

Ceramic sensors are fully interchangeable with hygrometers. Maintaining the measurement accuracy of the PURA hygrometer in applications where it is installed as a permanent on-line monitor is, therefore, simply a case of exchanging the sensor for a freshly calibrated one, which will be supplied with a traceable certificate of calibration. Sensor interchangeability is not affected by sensor cable length; therefore, this method of maintaining calibration can be used for all sensor installations.

For applications where the PURA hygrometer is not required for continuous operation, re-calibration of PURA hygrometer can be achieved by return of the complete instrument

- monitor, sensor and interconnecting cable - to Kahn Instruments.

# APPENDIX 1

### Register settings accessible by digital communication

Register number		Function					Read only

- Alarm Status

- N/A

- Processed Data – Result 				

- Processed Data – Channel 1 (dewpoint)		

- Processed Data – Channel 2 (pressure)		

- Setpoint 1

- Setpoint 2

- Setpoint 3

- Setpoint 4

10 to 23 			N/A

- Scale Value – Result

- Scale Value – Channel 1

- Scale Value – Channel 2

- Offset Value – Result

- Offset Value – Channel 1

- Offset Value – Channel 2

30 to 33			N/A

- D/A Zero – Analogue O/P 1

- N/A

- D/A Full Scale – Analogue O/P 1

37 to 64			N/A

- Hysteresis - Setpoint 1

- Hysteresis - Setpoint 2

- Hysteresis - Setpoint 3

- Hysteresis - Setpoint 4

69 to 70			N/A

- Make Delay - Setpoint 1

- Make Delay - Setpoint 2

- Make Delay - Setpoint 3

- Make Delay - Setpoint 4

75 to 76			N/A

- Break Delay - Setpoint 1

- Break Delay - Setpoint 2

- Break Delay - Setpoint 3

- Break Delay - Setpoint 4

81 to 128			N/A

- Cal Mode

- Code 1

- Code 2

- Code 3

- Code 4

- Code 5

- Code 6

- Code 7

- Code 8

- to 141			Reserved

- Setpoint 1 Control Register

- Setpoint 2 Control Register

- Setpoint 3 Control Register

- Setpoint 4 Control Register

146 to 147			N/A

- Brightness

- Baud rate Settings

- Serial Address

151 to 152			N/A

153				Model Number				

154				Version Number				

155 to 192			N/A

- Delay Type – Setpoint 1

- Delay Type – Setpoint 2

- Delay Type – Setpoint 3

- Delay Type – Setpoint 4

# APPENDIX 2

### Setup Codes

CAL		Calibration modes for input and output.		DEFAULT VALUE = 052

| Digit | 1st digit (left most) Calibration Mode | 2nddigit Calibration function | 3rddigit Object for calibration |
| --- | --- | --- | --- |
| 0 | Calibration functions as per 2nd and 3rd digit | No function | No function |
| 1 | Set baud rate, parity and serial address | Manual calibration (channel as per 3rd digit) | Processed result (dewpoint, ppm(V) and ppb(V) |
| 2 | N/A | N/A | Pressure input |
| 3 | N/A | N/A | N/A |
| 4 | N/A | N/A | N/A |
| 5 | N/A | Manual adjust | N/A |
| 6 | N/A | Analog output scaling (analogue channel as per 3rd digit) | N/A |

CODE 1	Tendency Indication, Additional LED’s, Display data source, flashing, Decimal points, rounding						DEFAULT VALUE = 200

| Digit | 1st digit (left most) Additional LED’s | 2nddigit Display data source | 3rd digit 0 only (no function) |
| --- | --- | --- | --- |
| 0 | LED annunciators are always off | Processed data - result (dew point, ppm(V), ppb(V)) | N/A |
| 1 | LED annunciators are on when relays are de-energized | Processed data - channel 1 (dew point only) | N/A |
| 2 | LED annunciators are on when relays are energized | Processed data - pressure | N/A |

CODE 3	Serial mode and analog output source		DEFAULT VALUE = 000

CODE 4	Channel 1 Measurement task, sampling rate	DEFAULT VALUE = 000

| Digit | 1st digit (left most) Analog sample rate | 2nddigit Analog output 1 source | 3rd digit 0 only (no function) |
| --- | --- | --- | --- |
| 0 | Single channel (60 Hz) | Dewpoint in degrees C | N/A |
| 1 | Dual channel (60 Hz) | Dewpoint in degrees F | N/A |
| 2 | Single channel (50 Hz) | N/A | N/A |
| 3 | Dual channel (50 Hz) | N/A | N/A |

CODE 5         Channel 1 Processing 		DEFAULT VALUE = 100

CODE 6	Channel 2 Measurement task		    DEFAULT VALUE = 000

| Digit | 1st digit (left most) Measurement task | 2nddigit 0 only (no function) | 3rd digit 0 only (no function) |
| --- | --- | --- | --- |
| 0 | Pressure in psig | N/A | N/A |
| 1 | Pressure in barg | N/A | N/A |

CODE 7	Result processing					DEFAULT VALUE = 000

| Digit | 1st digit (left most) Measurement task | 2nddigit 0 only (no function) | 3rddigit Range for result |
| --- | --- | --- | --- |
| 0 | Dew point | N/A | 1 – 9999 |
| 1 | ppm(V) | N/A | 1 – 999.9 |
| 2 |  | N/A | 0.01 – 99.99 |
| 3 |  | N/A | 0.001 – 9.999 |
| 4 |  | N/A | 0.01 – 99.99 (ppb(V)) |

Setpoint control 1 – 4	Relay latching, relay setup, source DEFAULT VALUE = 000

(Setup/Down menu after setpoints)

| Digit | 1st digit (left most) Relay sense | 2nddigit Setpoint source | 3rddigit Setpoint function for SP1 – SP4 |
| --- | --- | --- | --- |
| 0 | Relay energized above setpoint value | Processed result data (dew point, ppm(V), ppb(V) ) | No function |
| 1 | Relay energized below setpoint value | Dew point only | Relay latched |
| 2 | N/A | Pressure | De-energized relay |
| 3 | N/A | Open sensor connection | N/A |
| 4 | N/A | Open / short circuit sensor | Relay off for open sensor connection |

| 5 | N/A | Open / short circuit thermistor | Relay on for open sensor connection |
| --- | --- | --- | --- |
| 6 | N/A | All sensor faults | Relay toggles at 1 Hz for sensor connection fault |
| 7 | N/A | N/A | Setup hysteresis, make/break delay and delay type |

# APPENDIX 3

### Error Messages

The table below represents the possible error messages that can be seen on the hygrometer monitor under different conditions.

| Error message | Possible cause | Action |
| --- | --- | --- |
| Under | 1. Measured dewpoint is below the measurement range (-120°C/-184°F) | Contact Kahn Instruments regarding your application |
|  | 2. Sensor is out of calibration or damaged | Contact Kahn Instruments for sensor repair or recalibration |
| Over | 1. Measured dewpoint is above the measurement range (-40°C/-40°F) | Contact Kahn Instruments regarding your application |
|  | 2. Sensor was in contact with liquid water | Dry the sensor out with a stream of dry gas |
|  | 3. Sensor is out of calibration or damaged | Contact Kahn Instruments for sensor repair or recalibration |
|  | 4. Sensor is contaminated | Contact Kahn Instruments for sensor repair and application assistance |
|  | 5. Sensing element failure | Contact Kahn Instruments for sensor repair |
| Open | 1. Bad electrical connection between sensor and monitor | Verify connectivity of the sensor cable |
|  | 2. Internal sensor failure | Contact Kahn Instruments for sensor repair |
|  | 3. Internal monitor failure | Contact Kahn Instruments for monitor repair |
| ErrE | Internal monitor temporary memory corruption due to unusually high power voltage spike | Perform the following monitor resetting procedure: 1. Turn the power OFF. 2. Press and hold the  and  buttons. 3. Turn the power ON 4. Release the  and  buttons. 5. Verify the monitor settings (alarm setpoints, output scaling, etc.) |